## Slide 1: Medical Imaging
Chien-Min Kao
Associate Professor, Radiology and Medical Physics
University of Chicago

## Slide 3
CM: confocal microscopy; MPM: multi-photon microscopy; LOT: laminar optical tomography; (M)FMT: (mesoscopic) Fluorescence molecular tomography; OCT: optical coherent tomography; (M)DOT: (mesoscopic) diffuse optical tomography; PAT: photoacoustic tomography
Depth of penetration [log]

## Slide 5
Anatomy Structure
3D2D
Radiation
Wilhelm Röntgen
First Novel Laureate in Physics (1901)
Discovery of X-Ray: 11/8/1895
First “Medical” Image: 12/23/1895
Digital Radiography
Computer-Aided
Diagnosis (CAD)
X-ray (~100keV)
Tissue Attenuation Coefficient
(Electron Density)

## Slide 6
X-Ray Computed Tomography (CT)
X-ray (~100keV)
Tissue Attenuation Coefficient (Electron Density)
Anatomy & Structure
2001
~1975
Medical CT spatial resolution: < 1 mm
Micro CT spatial resolution: ~ 1-10 mm
Alan M. Cormack and Sir Godfrey N. Hounsfield received the 1979 Nobel Prize in Physiology or Medicine for the development of computer assisted tomography (CT)
Alan M. Cormack
(1924-1998)
Sir Godfrey N. Hounsfield
(1919-2004)

## Slide 7
Live or dead brain?
Functional vs Anatomical Imaging

## Slide 8
Functional imaging
“light up” activities

## Slide 9
Nuclear Medicine Imaging
Tracer Kinetics & Distribution
+ Radiolabeled chemicals
= Function/Physiology
Planar & Single-
Photon ECT (SPECT))
FDG-PET
Lung Cancer
In-111 Prostate Cancer Imaging
Tc-99m HDP
Whole-Body
Bone Scan
Tl-201 Cardiac Functional Scan
FDG-PET
stress
rest
 molecular imaging

## Slide 10
Antoine Henri Becquerel (1852-1908)
Pierre Curie (1859-1906)
Marie Curie (1867-1934)
Nobel Prize in Physics, 1903: Antoine Henri Becquerel, Pierre Curie, Marie Curie for the discovery of spontaneous radioactivity
  Marie Sklodowska
  Curie
Nobel Prize in Chemistry, 1911 by the discovery of the elements radium and polonium
Frederic Joliot
Irene Joliot-Curie
Nobel Prize in Chemistry, 1935: for the discovery of stable elements could artificially produce radioactive elements.
George
de Hevesy
(1885-1966)
Nobel Prize in Chemistry, 1941, for the use of isotopes as tracers in the study of chemical processes

## Slide 11
75th Chicago Pile-1 (CP-1) Commemoration
Manhattan Project December 2, 1942 Chicago Pile - One (CP1)
Enrico Fermi (1901-1954) received
the 1938 Nobel Prize in Physics for
his demonstrations of the existence of
new radioactive elements produced by
neutron irradiation, and for his related
discovery of nuclear reactions brought
about by slow neutrons.
On December 2, 1942
Man Achieved Here
The First Self-Sustaining
Chain Reaction
And Thereby
Initiated the
Controlled Release
of Nuclear Energy

## Slide 12
Historic Evolution to
Current Research
Post-WWII (1945)
“Atoms for Peace” Program
1953, Argonne Cancer Research Hospital (ACRH)
-- Peaceful Use of Atomic Energy in Medicine and Biology (both Diagnosis & Therapy)
1974, Franklin McLean Memorial Research Institute (FMI)
-- PET/SPECT
2005, Functional & Molecular Imaging Core (FMI)
-- Expanded into CT, Ultrasound, Optical Imaging, Emerging Technologies, Multi-Modality
Quantitative & Integrative Multi-Modality Functional & Molecular Imaging (QIM-FMI)

## Slide 13
ACRH Brain Scanner
1962-63 The Birth of Modern Nuclear Medicine
First Tc-99m Brain Scan
“First Molecular Image”
Multi-Disciplinary ACRH Molecular Imaging Team
Paul Harper (Surgeon)
Robert Beck (Physicist)
Katherine Lathrop (Chemist)
Donald Charleston (Engineer)
Alex Gottschalk( Radiologist)
World’s First Tc-99m Brain Image, 1963
New Disciplines at Interfaces of
Biology, Medicine, Physics, Chemistry, Mathematics, Computer/Computing Science, Material Science/Engineering, Electrical Engineering + X

## Slide 14
p  n + e+ + n + energy
Carl David Anderson (1905­-91)
received the Nobel Prize in 1936 for the discovery of the positron.
At age 31, Anderson was then the youngest person to receive the Nobel Prize.
Positron Emission Tomography (PET): Principle
E = mc2

## Slide 15: Positron-Emitting Nuclides

## Slide 16: Production of Isotopes (Cyclotron)
Modern Medical Cyclotron
RadioChem
Synthesis
Module
UChicago new cyclotron and radiochemistry facilities (2017)

## Slide 17: Production of Isotopes (Cyclotron)
18O (p,n) 18F
At the ion source
of the 184-inch
cyclotron in 1948.
The first cyclotron
is built in late 1930
Lawrence at the
Controls of his
cyclotron in Berkeley.
Ernest Orlando Lawrence (1901-1958)
received the 1939 Nobel Prize in physics
for the invention and development of
the cyclotron
CS-15 Installed at UChicago-ACRH/FMI in 1968

## Slide 18
Positron Emission Tomography (PET): Principle
Coincidence Detection
LOR, li
LOR sensitivity
(calibration)
subject attenuation
(measured using a tx source, or calculated from CT, MR images)
line integral
/ray sum

## Slide 19
Image Reconstruction from projections:
2D FBP Demonstration
Positron Emission Tomography (PET): Principle

## Slide 20
PET system
Share similar designs with HEP detectors but subject to different size and cost constraints and performance requirements

## Slide 21: Scintillation Detector
gamma rays
scintillator
electrical signal
  time
  energy
Photomultiplier (PMT)
photocathode
106 e/ph
e-
anode
dynode

## Slide 22: PET Event Detection
Energy window (400-650 keV)
Qualified energy?
Qualified energy?
Coin?
yes
yes
yes
Crystal IDs
(LOR)
Coincidence timing window (4-10 ns)

## Slide 23: PET Event Types
Scatter (B)
  Compton scattering in subject
  energy <511 keV
  depend on object size and scanner geometry
  distribution affected by subject shape
  cannot be easily measured
Randoms (C)
  accidental coincidences
  smaller coincidence windows  fewer randoms depend on activity levels
  relatively uniform distribution
  can be measured using delayed coincidences

## Slide 24: PET Block Detector
(https://www.radiologycafe.com/radiology-trainees/frcr-physics-notes/pet-imaging)

## Slide 25: Factors affecting resolution
FWHM = d/2
d
Crystal size
Positron range
depends on Emax of emitted positrons
              18F rms = 0.23 mm
              11C rms = 0.39 mm
range is inversely proportional to the absorber density
Photon acolinearity
±0.25°
Intercrystal scattering
significant for small crystals

## Slide 26: Block Detectors
A
B
D
C
Event energy (pulse height) and time are derived from the summed pulse, assuming only one event within the processing time
Compromised count-rate capability
More lights (brighter scintillators and better light collection)  better statistics  better energy resolution, better spatial resolution, better timing
Faster scintillator and PMTs  faster timing
Clinical PET: 4 mm detector pixels, ~80 cm diameter, 20 cm length  31,400 pixels
8x8 block detectors  reduce electronic channels by x16
G Tarantola (2014)

## Slide 27
(R. Lecomte, NIM 2004)
FWHM= kR C2 + (d/2)2 + R2 + B2

## Slide 29: Depth-of-interaction (DOI) blurring
(from Siemens Medical)

## Slide 30: Large Patient Problem
58 kg 89 kg 127 kg
(J Karp et al)

## Slide 31: Time-of-flight (TOF) PET
Non TOF
TOF
equal scan time
(http://flc.desy.de/pet/intro/index_eng.html)
(M Conti, EJMMI Physics, 2011)

## Slide 32
PET/CT
~500-600 ps with LYSO/PMT

## Slide 33
Improve TOF resolution

## Slide 34: Silicon Photomultiplier (SiPM)
SiPM/SPM/MPPC array
SiPM/SPM/MPPC
1-6 mm pixel, array format, 30-70 V, robust, MR compatible, TOF capable, low-cost potential
~1 inch, >1000 V, fragile, very sensitive to magnetic fields, TOF capable

## Slide 36: Compact SiPM based PET detectors
double-ended
Side readout
(SR Cherry et al)
(CS Levin et al)
Multi layered
(A. Del Guerra et al)
Can support TOF also

## Slide 37
PET/MR
Light Tight RF Shield with copper coating
LBS array with (ESR)
SiPM with circuit boards/ASICS
Thermal coupling
Aluminum mounting
25 cm AFOV
10.5% ER
390 ps CRT

## Slide 38
Design Approaches

## Slide 39
SiPM Based Inserts
16 detector modules
Axial FOV = 55m, ring diameter = 64mm
Carbon fiber tube for shielding
ID/OD: 60mm/100mm
1.2x1.2x10mm LYSO, 9x9 arrays, pitch 1.28mm
4x4 Hamamatsu SiPM array
Resistor charge network reduces output to 4
Resolution (OSEM/FBP): 0.75/1.31mm at center, 1.46/2.18mm at 14mm off center
CRT ~ 1.33 ns
Temperature and voltage controls
(GB Ko et al, Med Phys 2015)

## Slide 40
SiPM

## Slide 41
SiPM Based Inserts
1.2x1.2x10mm LYSO, 9x9 arrays, pitch 1.28mm
4x4 Hamamatsu SiPM array
Resistor charge network reduces output to 4
(GB Ko et al, Med Phys 2015)

## Slide 42
Stripline (delay-line) Readout
PD outputs
Digital Signal Processing (FPGA)
high-speed serial output
SLs
DAQ
DAQ
DAQ
DAQ
DAQ
DAQ
PDs
preamp
The position of firing PD is determined using the propagation time difference.

## Slide 43
UChicago detector module
2.54 cm
Power connector
LYSO/SiPM arrays: 4x8 pixels (MPPC), 3.2mm pitch, 10mm thick
curvy SLs to increase separation between SiPMs to 25 mm and 400 ps
~2.62 cm
~5.0 cm
SL outputs x8
excellent pixel discrimination
Best CTR ~ 250ps
LED ~3 mV

## Slide 44
Stripline Readout is Scalable
16 SiPMs/Stripline
32 SiPMs/Stripline

## Slide 45
FPGA-only MVT sampling DAQ
Compact, high channel density
FPGA based: affordable, flexible, upgradable
Digital signal samples and serial digital output: robust, universal, digital signal processing
Clean IO interface and network ready: rapid system development
Do not require any proprietary parts and costly and length ASIC development
Intrinsic time resolution ~ 300ps
MVT FPGAs
Ethernet
amplifiers
DAC
72ch board
Control FPGA
Sync clock
power

## Slide 46
PET Insert
16 pxiels per stripline
14 modules; 2.5 cm axial FOV
Initial phantom image
(3.0, 2.6, 2.2, 1.8, 1.4, 1.0 mm)

## Slide 47
Stripline Readout/MVT DAQ
Substantial channel reduction
Alleviating engineering issues: heat dissipation, power consumption, space constraint.
Cost saving
Detached electronics
  Compactness
  mitigate RF and temperature control issues in PET/MR
Highly flexible, scalable DAQ electronics
With SL readout, multiple detectors can be daisy chained
The DAQ provides high channel density and is affordable
The DAQ produces digital samples
Use commodity electronics components and industry I/O standards

## Slide 48
Bruker Peclinical PET Insert
eight 50x50x10mm continuous detector blocks, 12x12 SiPMs (SensL), resistive network row/column readout
High resolution w/ DOI
ID/OD = 114/198 mm
Carbon Fiber RF shielding
~ 1m long
(AJ González, NSSMIC 2016)

## Slide 49
Tilable Detector Block/Panel Detector
16x16 LYSO
16x16 SiPMs
SL board
SL board
~50x50mm detector, 4 outputs per detector, DOI and TOF capable
Power
Sync generator
Network switch
PC/Workstation
High-speed serial data, power, sync-clock cables

## Slide 50
Other trends: Dedicated Brain PET
PhotoDiagnostic Systems (PDS) NeuroPET/CT
Siemens HRRT
Need higher resolution (1-2 mm) and sensitivity for Brain imaging
Mobile units can be useful for emergence care (stroke evaluation) and surgical room
UChicago-HUST

## Slide 51
Wearable Brain Imagers
(CE Bauer, Brain Behavior 2016)
To optimize sensitivity and image quality

## Slide 52
Other Organ-Specific Systems
PEM Naviscan
Newsoft
Cardiac
PET
Notes: 241

## Slide 53
Probes

## Slide 54
EXPLORER: Total Body PET
From http://explorer.ucdavis.edu/

## Slide 55
Proton therapy/In-Beam PET
T. Nishio et al.: Med. Phys. 33 (2006) 4190